## Slide 1: BAMRR MRI Safety Week 2024
What happened- the incident
What is a significant adverse event
How to share learning
What happened afterwards
Recommendations and follow up
What is next in ‘incident reporting’ world

## Slide 2: BAMRR MRI Safety Week 2024
https://policyonline.nhslothian.scot/wpcontent/uploads/2023/03/Adverse%20Event%20Management%20Procedure.pdf
Flowchart of actions to be taken to effectively manage adverse events

## Slide 3: BAMRR MRI Safety Week 2024
An investigation was undertaken using a multi-disciplinary team and contributory factors were identified using a human factors approach.
Recommendations were made using a ‘Systems Engineering Initiative for Patient Safety’.
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary Edinburgh

## Slide 4: BAMRR MRI Safety Week 2024
The investigation focussed on the following factors, which will be discussed in turn:
2.
Patient
3.
Staff
1.
Task
&
Technology
6.
Organisation
& Management
7. Institutional
Context
4.
Team
5. Work Environment
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary Edinburgh

## Slide 5: BAMRR MRI Safety Week 2024
Departmental Layout:
Equipment Rooms
3T Scanner Room
Control Room
Induction Room
Recovery Room
Trolley bay
CT Scanner Room
wc
wc
Changing
room
Changing
room
Adjacent to DCN
MR CAA
The MR CAA is delineated by the orange lines.
View of MR CAA
from main entrance
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
Unfortunately, there are multiple access points to the MR CAA The main entrance is through this door but there are additional access points via the induction room, recovery room and through CT scanner room
- - - - -

## Slide 6: BAMRR MRI Safety Week 2024
3T Scanner Room
Control Room
Induction Room
Recovery Room
Trolley bay
1
2
3
4
5
1.
Task
&
Technology
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
Finding:
Process Failings

## Slide 7: BAMRR MRI Safety Week 2024
Finding:
Reduced Focus on MRI
3T MRI
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
2.
Patient

## Slide 8: BAMRR MRI Safety Week 2024
Finding:
Distractions to the Main Task
3T MRI
Radiographers – control room
Anaesthetic Consultant
- Induction room
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
3.
Staff

## Slide 9: BAMRR MRI Safety Week 2024
MR Radiographers
Anaesthetists
Operating Department Practitioner
Radiology Department Assistant
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
4.
Team
Finding:
Lack of Clarity in Roles and Responsibilities

## Slide 10: BAMRR MRI Safety Week 2024
3T MRI
Control Room
Induction Room
Recovery Room
CT
Finding:
Multiple staff groups waiting in MR CAA
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
5. Work Environment

## Slide 11: BAMRR MRI Safety Week 2024
MR Safety training update due!
Standard Operating Procedure
Finding:
No written SOP’s and Training records not up to date
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
6.
Organisation
& Management

## Slide 12: BAMRR MRI Safety Week 2024
Finding:
Royal Hospital for Children and Young People
had recently moved to a new building
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
7. Institutional
Context

## Slide 13: BAMRR MRI Safety Week 2024
Slide courtesy of
Dr Gillian MacNaught
Lead MRI Physicist
Royal Infirmary
Edinburgh
Putting all these investigation findings into a timeline for the incident, this shows the series of events that led up to the critical moment, resembling the well known ‘Swiss cheese’ model of incidents when all the gaps or holes align…

## Slide 14: BAMRR MRI Safety Week 2024
BAMRR are also pleased to direct you to what BIR and ISMRT and SCoR have produced for MR Safety Week
https://bir.org.uk/get-involved/special-interest-groups/bir-magnetic-resonance.aspx
https://www.ismrm.org/mr-safety-links/mr-safety-week-2024/#:~:text=These%20virtual%20meetings%20are%20free,%E2%84%A2%20and%20Category%20A%20Credit.
https://www.sor.org/
Materials from MR Safety Week from previous years are available on the BAMRR website https://www.bamrr.org/safety/